# Pyloriset Dry

Rapid Detection of Helicobacter Pyloi antibodies

Orion Diagnostica´´s Pyloriset Dry is a latex agglutination test for the rapid detection ofHelicobacter pylori antibodies in serum. Pyloriset Dry is an easy and inexpensive mean for preliminary screening of patients prior to endoscopy. No special instrument are required, and the test result are available within minutes.

The performance of Pyloriset Dry is shown below. Study 1 shows the results obtain when 90 randomly selected patient samples were investigated by culture, histology, rapid urease test and Phyloriset Dry . Study 2 shows the results obtained when samples taken from 123 symptomatic children were tested using Pyloriset Dry, histology and / or culture.

|  | Sensitivity | Specificity | PPV | NPV |
| --- | --- | --- | --- | --- |
| STUDY 1 | 93,30% | 95,60% | 95,50% | 93,30% |
| STUDY 2 | 85,50% | 98,10% | 98,30% | 84,10% |

Test Procedure

Allow the reagents to reach room temperature before use. Note: There are two alternative methods for diluting the serum samples (1A or 1B).

1A . Dilute the serum sample 1:4 with the dilution buffer in a test tube.

Continue as in point 2.

1B. The serum can also be diluted directly on the test card. Pipette 30 ul of the dilution buffer next to the latex reagent. Pipette 10 ul serum onto the dilution buffer drop. Mix thoroughly. Continue as in point 3.

2 . Pipette 40 ul of the diluted serum sample onto the circle of the test card next to the latex reagent.

3. Mix the diluted serum samples with the latex reagent on the card.

4. Tilt and rotate the test slide. The result is available within 3 minutes.

5 .Test result. Positive result: Circle to the right and in the middle.

Negative result: circle to the left.

Pyloriset Dry

Contents of the kit 24 test		Cat. No. 68194

Test cards with dried latex reagent  4x2 cards containing H. Pylori antigen

Positive control 			0.5ml

Negative control			0.5 ml

Dilution buffer (PBS, Ph 7.2)	30 ml

Mixing sticks			30 sticks

Plastic pouch

Instruction for use

Storage

Store the Pyloriset Dry test kits in a refrigerator at +2 … +8º C.

Do not freeze. The expiry date is marked on the box.

Pyloriset seco

La detección rápida de anticuerpos Helicobacter Pyloi

Diagnostica' Orión Pyloriset seco es una prueba de aglutinación en látex para la detección de anticuerpos rápida ofHelicobacter pylori en el suero. Pyloriset seco es un medio fácil y barato para la selección preliminar de los pacientes antes de la endoscopia.No se requiere de instrumentos especiales, y el resultado de la prueba están disponibles en cuestión de minutos.
El rendimiento de Pyloriset seco se muestra a continuación.Estudio 1 muestra los resultados se obtienen al azar 90 muestras de los pacientes fueron investigados por la cultura, la histología, test rápido de ureasa y Phyloriset seco. Estudio 2 muestra los resultados obtenidos en las muestras tomadas de 123 niños sintomáticos se sometieron a pruebas Pyloriset seco, histología y / o la cultura.

 Sensibilidad y especificidad VPP VPN
ESTUDIO 1 93,30% 95,60% 95,50% 93,30%
ESTUDIO 2 85,50% 98,10% 98,30% 84,10%

Procedimiento de prueba

Dejar que los reactivos alcancen la temperatura ambiente antes de su uso. Nota: Hay dos métodos alternativos para diluir las muestras de suero (1A ó 1B).

1A. Diluir el suero de la muestra 1:4 con el tampón de dilución en un tubo de ensayo.
Continuar como en el punto 2.

1B. El suero también puede diluirse directamente en la tarjeta de prueba. Pipeta 30 ul de tampón de dilución al lado del reactivo de látex. Pipeta 10 ul de suero en la caída de tampón de dilución.Mezclar bien. Continuar como en el punto 3.

2. Pipeta 40 ul de la muestra de suero diluida dentro del círculo de la tarjeta de la siguiente prueba con el reactivo de látex.

3. Mezclar las muestras de suero se diluye con el reactivo de látex en la tarjeta.

4. Inclinar y girar el portaobjetos. El resultado está disponible en 3 minutos.

5. Resultado de la prueba. Resultado positivo: Círculo a la derecha y en el medio.
                      Resultado negativo: el círculo a la izquierda.


Pyloriset seco
Contenido del kit de prueba de 24 gatos. N º 68194
Tarjetas de prueba con el reactivo de látex seca tarjetas de 4x2 que contiene el antígeno de H. pylori
0.5ml control positivo
Control negativo de 0,5 ml
Tampón diluyente (PBS, pH 7.2) 30 ml
Palitos de mezcla de 30 palos
Bolsa de plástico
Instrucciones de uso

Almacenamiento

Guarde el estuche Pyloriset seco de prueba en un refrigerador a 2 ... 8 ° C.
No congele. La fecha de caducidad marcada en la caja.